## Slide 1
27 January 2023: Cable under test with Laser Numero 3-2
DP2
1m fibre
After 1m fibre:
I= 30A => P1= 37.5 m W
I= 34A => P2= 47 m W
coupling
2
DP2
1
Direct power output from DP2-1:
I= 30A => P1= 62 m W
I= 34A => P2= 79 m W
DP2
1m fibre
After fibre 1 to 6 of cable N3-2
coupling
3
Test fibre cable 3-2
# fibre | I (A) | P (mW) | Full Attenuation (dB) | Fibre loss Estimation (dB)
1 | 30 34 | 14 17 | 4.4 | 2.9
2 | 30 34 | 14 17.5 | 4.2 | 2.7
3 | 30 34 | 13.2 16.5 | 4.5 | 3
4 | 30 34 | 13.6 17 | 4.4 | 2.9
5 | 30 34 | 13.3 16.5 | 4.5 | 3
6 | 30 34 | 13.2 16.6 | 4.5 | 3
Attenuation between 2) and 1), adding 1m fibre and coupling:
IL(dB) = 10log 47/79 = 2dB
=> FC connector is typically 1 or 1.5dB :
Fibre loss Estimation = Full attenuation – 1.5
Fibre loss according to OFS: 13dB/km:
150m of fibre should be 1.95 dB
Fibre loss according to Liyuan based few years ago: 25dB/km
150m of fibre should be 3.75 dB
The 25 dB/km was measured for 440 nm with Ti:Sapphire laser as shown on page page 8 in Ren-Yuan's talk (http://www.hep.caltech.edu/~zhu/talks/redlaser_020416.pdf). I see 488 nm and 532 nm are listed with 13 dB/km on the web, while the loss of 532 nm is basically consistent with our measurement, the 13 dB/km at 488 nm is a few dB smaller than our measurement, OFS may have made some improvement. Hopefully, your measurement can confirm this improvement at
447 nm.
Liyuan

## Slide 2
2 February 2023: test 2 more cables
# fibre | Cable 3-2 P (mW) @34A | Full Attenuation (dB)
1 | 17 | 4.4
2 | 17.5 | 4.2
3 | 16.5 | 4.5
4 | 17 | 4.4
5 | 16.5 | 4.5
6 | 16.6 | 4.5
# fibre | Cable 2-3 P (mW) @34A | Cable 2-3 Full Attenuation (dB) | Cable 3-3 P (mW) @34A | Cable 3-3 Full Attenuation (dB)
1 | 16 | 4.5 | 16 | 4.5
2 | 16.5 | 4.4 | 16.6 | 4.4
3 | 15.8 | 4.6 | 15.9 | 4.5
4 | 16 | 4.5 | 16 | 4.5
5 | 17.2 | 4.2 | 16.5 | 4.4
6 | 16 | 4.5 | 16.5 | 4.4
After 1m fibre:
I= 34A => P2= 45.5 m W
After 1m fibre:
I= 34A => P2= 47 m W
Full Attenuation = 10 log P/47
Fibre loss Estimation = Full attenuation – 1.5
Full Attenuation = 10 log P/45.5
Fibre loss Estimation = Full attenuation – 1.5
Attenuation for 3 cables are identical, no difference for the fibre #5