## Slide 1
David BAILLEUX
On behalf of the Caltech group
Issues for the Laser Monitoring System M&O
- US CMS ECAL IB meeting 19 March 2009 -
Chilled water supply
Residual B field effect
Laser system DCS
Spare parts and upgrade

## Slide 2
Laser System M&O
Chilled Water Supply
 Status during July 2008 cooling meeting:
Pipe work completed; work started on pumping station;
Electrical control box ordered.
 Pumping station tested in February, 2009, for one laser:
- Only one pump is working (another piece in reparation);
Low pressure alarm not working: pump still ON  leak detection not working;
Missing bypass as it was installed in H4  else the pump stopped if there is no water flow;
DCS not tested (e-mail notification etc.)
Missing electrical connection for barracks control.
Up to now:
Mixed water used for lasers cooling (racks) : dirty water, no leak detector.
Water leak may happen as in the past: consequence for all racks ? Responsibility?
Is laser can still work the night and weekend like this ?
Task follow by TS/CV : not schedule yet. Werner and Wolfgang are aware of that.

## Slide 3
Residual B Field Effect
One new digital infrared viewer ordered and received .….…
One digital scope with LCD display in hand ……………………….
One power meter ordered and received ……………………….….
One AC control unit has to be moved inside the room …….
Shielding around power supply and pump laser (lamp) ……
Interlock on laser lamp when B is on fast variation .………….
Laser System M&O
A meeting on February 12 approved following solutions:
Will be tested in June when B filed resumes. If successful  2 more DSO and instruments will be ordered
Planned: waiting..
To be planned. Wolfgang has to put this on agenda.

## Slide 4
DCS for the Laser System
Priority need to be defined:
 Room temperatures: CANbus installed, waiting connection and read out via central DCS.
 Laser settings: Laser settings should be online and recorded DCS/PVSS (current, delay, ect.)
On the laser side, the information is on the DIM server so ready to be collected.
Wolfgang is aware of this.
Laser System M&O
Optical switch: no spare 1 x 100 switch (old one is 1 x 80 so could not covered EB and EE)
Pump laser upgrade: The YLF:Nd pump laser may be replaced by diode pumped model with high stability, reliability and much less maintenance requirement.
Spare Parts and Upgrade