## Slide 1: Discussion today
Quantronix option has an unacceptable delivery time (150-180 days ARO)
  Excluded from further discussions
Need to decide between two options from Photonics:
  440nm Ti:S + diode-pumped laser (this is the baseline option endorsed already by US-CMS project management)
  447nm diode-pumped laser (new option since Renyuan’s visit to Photonics on October 19th)
If we decide upon the 447nm option we need to seek endorsement from US-CMS PM before proceeding with the order
  And we would start this immediately after this meeting

## Slide 2: How do we decide?
Stability is the critical issue here
  Short-term (days/weeks)
    Governed by laser parameters such as:
      pulse width
      jitter
      pulse-to-pulse amplitude/width variation
  Long term
    Slow drift of above quantities
    Frequency of maintenance interventions (requiring re-tuning etc.)

## Slide 3: Technical specs & dimensions
 | Photonics 440 | Photonics 447
Frequency | 440nm | 447nm
Pulse energy @ 100Hz | 1mJ | 0.9mJ @ 85us delay
Pulse width | <25ns @ 100 Hz | <15ns @ 0.9mJ
Pulse jitter | <3ns RMS | <3ns RMS
Delay | <5us | 85us for 0.9mJ
Repetition rate | 0100 Hz | 0200 Hz
Size of laser unit | 368 x 660 x 127mm | 190 x 560 x 95mm
Power supply unit | 483 x 476 x 178mm | 343 x 483 x 133mm
Water Cooling unit | Required | Not needed
Delivery time | 12-16 weeks | 12-14 weeks
Comments:
1) both options satisfy our technical specifications
2) Q-switches (Pockel Cells) are included in both lasers, to obtain the 3ns jitter
3) the repetition rate offered by the P447 could be useful….
4) the P447 is far more compact than the P440 and requires no external cooling
5) spare parts for both options can be purchased in advance and are readily available

## Slide 4: Pros and cons of 447nm option
Pros
  Considerably simpler and more compact than 440nm option
    No Ti:S system
    No external cooling system required
    Fewer parts that could fail
  Low power consumption
  Testimonial from Zebra Imaging Inc., using laser in 24/7 mode for ~1.5 years very positive
Cons
  Not the same wavelength as existing blue lasers
    Should not be a problem
  Significant delay (90msec) required to achieve 1mJ output
    90msec delay requires modification of L1 firmware (and subsequent change of sw/fw of existing hardware to use same delay)
    Could operate at, for example, 85msec and 0.9mJ, without changing firmware
  Diode operated in pulsed mode to achieve high peak power (due to weak 1342nm line). Pulse shape may be less stable than in DC mode
  Proprietary technology with few similar lasers in existence

## Slide 5: Photonics Testimonials #1
General
  Dr. Yuan Liu of Oakridge National Lab commented on the Nd:YLF pumped Ti:Sapphire laser system from Photonics Industries procured about one and half year ago. The laser system is run at 10 KHz in 24/7 mode for several weeks for each run. While the Ti:Sapphire laser has no problem, the Nd:YLF had a problem caused by condensed water and it was fixed by Photonics Industries. She recommended Photonics Industries.
  Dr. Balakrishnan of University of Washington commented on their Nd:YLF pumped OPO system from Photonics Industries purchased about 5 yrs ago. They are basically satisfied with the laser. The laser is run at 1 KHz, but NOT in 24/7 mode. The original diode module is still in good shape with accumulated time of over 3,500 hrs. There were some small issues like power dropping, chiller not working properly etc. The service is not as good as expected. While hoping Photonics will improve its service, they recommended the Photonics Industries.

## Slide 6: Photonics Testimonials #2
447nm
  Liyuan has talked to Michael A. Klug, CTO of Zebra Imaging Inc. (ZI), who bought two 447 nm lasers from Photonics Industries (PI). Mr. Klug's comments are as follows.
  ZI procured two 447 nm lasers from PI about 1.5 years ago. The lasers are run in 24/7 mode at 120 Hz with pulse energy of about 1 mJ and pulse width of about 15 ns. The required instability of laser pulse energy is about 3% rms and jitter < 3 ns rms. There is no requirement for the delay from external trigger for their application. While do not have laser diagnostic data registered, they believe that the long term stability is rather good by looking at the imaging threshold. These two 447 nm lasers have been run for about 7,000 hours and no significant degradation is noticed.
  ZI started using PI lasers more than 5 years ago. They have about 15 PI's lasers, most of them are green (532 nm). The average pump diode lifetime is about 15,000 hrs. They are satisfied by the PI's lasers.