## Slide 1
Lasers Incidents history
8-9 April, Easter: first: http://cmsonline.cern.ch/cms-elog/759922
Laser1: unstable pulse. Current change from 21 to 25 A by shifters.
Laser2 : in good shape. Control OK. Only lake of experience to put it back ONLINE. It was not possible to do it remotely because main fibre was remove and TiS box in manual mode from last week to test the new optical switch with Guy, BUT this was not a reason why it was not possible to put it online.
Photonics: several issue which prevent normal operation:
GPIB failure. This is most probably due to ICS box. Reboot it solved the problem.
Occurred Monday and this night 23h20 and maybe this weekend too. It’s appeared after ~5hrs of data taking.
Voltage supply change yesterday. I will change the box itself with new one today to see.
Shutter: yesterday found communication problem with the new fast shutter.
 shutter OK on ecal-laser-room-08, laser control of DP2-447
 shutter NOT OK remotely through the MAIN laser control ecal-laser-room-02. Shutter DP2 error, not moving
Guy told me this was the same problem this weekend.
I found this problem: PC name ecal-laser-room-02 change locally to ecal-laser-room-06 automatically, old name of this PC. (Ecal-laser-room-06 unknown on the web). This issue has been solved yesterday (Pedro)
. Friday 7 was the first time of long data taking with 3 lasers . Thus this is a kind of commissioning , and normal to find some bug and not easy to investigate because of new system
. Non experts can give advice but at the end to many opposite or wrong information which make a mess at the end mainly for new on call shifters.

## Slide 2
Lasers Incidents history
29 Avril- 2May:
1) GPIB error Sunday 29, 21h28. Restart all 00:57 am.
Comment: linearity scan for the ENCAP finished at ~21h10.
2) Laser1 OFF- restart: 01 MAY, 19h10.
Restart of laser1 OK and sequence restarted ~20h20. Laser1 is ‘fault state’. I don’t know what was the error display on the laser...
Comment: - linearity scan start at 17h08 , stop at 18h52
- power glitch at 6am, following by CMS wide water cooling disturbance
3) Laser1 OFF  Laser2: 02 MAY, 2h54
Laser2 ON at 5h20.
4) Laser2  Laser1 : 02 MAY, 16h30.
Laser1 investigation and problem:
No problem at all regarding the operation: device rebooting for 1) and spare laser2 online OK thus all OK regarding ecal laser operation at it is defined. Good !
The others things are details and transparent regarding ECAL.
Laser1 investigation:
- Put back 20 degrees on room temperature at is was to avoid light disparity (was put to 18°)
- Laser1 error : “water flow error”. While Emanuele got “water Di error” . The error remain ON all the time despite resetting the laser power supply (and error before the pump start)
- Swap the 2 chips relay on logic unit: no effect
- Replace the flow sensor with spare one: no effect
- Replace one relay chip on logic unit: OK , error disappeared !
- Start laser: lamp can’t start ignition , no current à dismount it. Found lamp broken. Succeed to remove piece of glass inside the flow tube without dismount all.
- Restart laser with no tuning at ALL to avoid big monitoring step. Power of TiS laser same as 13 April (new lamp)

## Slide 3
Lasers Incidents history
Conclusion :
i) I m not sure why laser1 went OFF in 2). which error ? If it was “water temperature error” that’s mean no input cooling, due to CMS water trouble -> Most probably the case because the 1 May laser1 was not used so no visible defect before the afternoon when ECAL start to use it.
Anyway laser restart after that thus nothing to suspect on laser side.
BUT this is for sure very bad for the lamp in general when the laser stop suddenly. This for me can explain why the lamp brake later on, also because laser1 not only stopped once but twice like this .
ii) For the case 3), failure of a chip happened already in the past 2 or 3 times (thus I have spare). Maybe there was disturbance on the power line from the laser which make this electronic chip in bad shape after a while and then failed later on ! this is not excluded but this is not important: nothing to be done regarding power line and laser electronic for a such rare default.
iv) The following 3 things which was mentioned have nothing to do with laser default:
- Level of water
- Temperature of water
- Room temperature

## Slide 4
Lasers Incidents history
8 MAY
Green laser out of the sequence ~ 4:00pm http://cmsonline.cern.ch/cms-elog/764979
Guy’s investigation remotely : problem with VNC and its terminal crashed something like this. Emanuele discovered the PC02 was logged OUT thus no more laser DAQ that’s explain no more laser during this time. ~6:30pm
http://cmsonline.cern.ch/cms-elog/764988
Recover laser DAQ with help of Kejun before there were error restarting program 8:20pm
Origin of Green laser problem unknown and this is the task of Saclay’s group (we have no control on it)
The restart of laser DAQ make some gpib conflicts but problem may not appeared without intervention.
9 MAY 4am:
no more DP2 on the sequence discovered by myself at 7am. Restart only the GPIB driver script which solved the problem (GPIB_CONFIG)
http://cmsonline.cern.ch/cms-elog/765053
9 MAY 4pm:
no more DP2 on the sequence, called by DOC: restart the ICS box GPIB this time.
http://cmsonline.cern.ch/cms-elog/765089
Action:
-> reboot PC during ecal run the 11 May to see how to recover and restart DAQ , for laser shifters also. Clear instruction.
-> program will be updated to restart the GPIB_CONFIG automatically if it’s crashed: 1 of the 2 problem may disappeared.
-> GPIB problem: ICS box to be rebooted, no other way right now.

## Slide 5
Lasers Incidents history