## Slide 1
Lasers inspection with Liyuan, 29March
1) DP2-447 lasers output: no evident degradation is observed. The 1st DP2-447 has a total runtime of 15K hrs (about 2 times of the warranty time), indicating that running the laser at 35 A pump current (with a maximum of 50 A) increases the lift time of the pump diodes. The pulse energy, width and jitter are evaluated with internal trigger and a short term run of 1~2 hrs, they are all well within the specification.
2) The ancillary optics of the two DP2-447 lasers are checked and the total transmittance is measured, no damage and degradation are observed. The output fiber of the 1st DP2-447 laser is polished and cleaned, and the coupling alignment is optimized. While the power measured after the 1x100 optical switch is increased by about 10%, the monitoring pulse amplitude at the fast monitor and LightChecker doesn’t linearly increase, which may be caused by mismatch of the area of PIN diode with the beam size or misalignment of the beam on the PIN diode in the monitor box. More diagnostic is needed to figure it out. If it’s true, we may insert a beam matching lens in front of the PIN diode to improve this power monitoring linearity in near future.
3) The Quantronix Ti:S IR laser is also checked and tested, it’s basically in working condition, but long term stability may be a concern. We studied IR laser choices more than a decade ago, and may proceed to investigate possibilities in the market if it is needed for ECAL monitoring upgrade.
Before intervention
After intervention

## Slide 2
Maintenance and calibration done in February already during YETS: